Assembly Instructions: Manawa Ora Ventilator

Please Refer to drawings MO-GA001, MO-PID-001, MO-E-001, MO-E-002 for assembly detail. Steps below outline an overview of the assembly process.
1. Assemble pipe fittings and solenoid valves for inspiratory line as per MO-GA001. Take care to ensure thread tape is used on all male threads. Tighten fittings hand tight and then a further 3/4 turn or more using a spanner.
2. Assemble pipe fittings and solenoid valve for expiratory line as per MO-GA001. Take care to ensure thread tape is used on all male threads. Tighten fittings hand tight and then a further 3/4 turn or more using a spanner.
3. Lower expiratory line into housing. Mount tank fittings onto housing and expiratory line as per MO-GA001.
4. Lower inspiratory line into housing. Mount Tank fittings onto housing and inspiratory line as per MO-GA001.
5. Mount all sensors in T-junctions as per MO-GA001 & MO-PID-001. Ensure sensors are mounted using correctly modified 1” BSP blanking plugs.
6. Mount ‘raspberry pi controller’, ‘12V Relay Board’, ‘12-5V DC-DC’ converter and ‘230V AC 12V DC UPS’ as per MO-E-001
7. Mount the above electrical controls into ventilator housing as per MO-GA001

1. Connect solenoid valves and flow sensors to 12V Relay Board as per MO-E-001 & MO-PID-001
2. Connect humidity and pressure sensors to 5V rail on Raspberry Pi Controller as per MO-E-002 & MO-PID-001
3. Once device has been fully assembled. Consult MO-ITP-001 for Factory acceptance testing