## Sheet1
Prices in NZD
Hardware |  | Medical grade |  |  | Cheap grade |  |  | Washing machine version version
|  | Example | Unit price | Price | Example | Unit price | Price | Example | Unit price | Price
Controller | 1 |  |  | 0 | Arduino | 40 | 40 | Arduino Uno or similar | 40 | 40
Standard hospital pneumatic fittings | 2 |  |  | 0 |  |  | 0 | air compressor fittings | 100 | 200
Manual pressure regulator | 3 |  |  | 0 |  |  | 0 | LPG adjustable regulators (6"-18" W.C) | 30 | 90
Solenoid valve | 3 |  |  | 0 |  |  | 0 | washing machine solenoids | 80 | 240
O2 meter (flow, concentration, pressure) | 2 |  |  | 0 |  |  | 0 | Car o2 meter or electrochemical o2 meter | 100 | 200
Inline HEPA filter | 1 |  |  | 0 |  |  | 0 | standard fuel filter housing with HEPA filter material | 5 | 5
Exhaust HEPA filter | 1 |  |  | 0 |  |  | 0 | standard fuel filter housing with HEPA filter material | 5 | 5
Box | 1 |  |  | 0 |  |  | 0 |  |  | 0
Batttery | 1 |  |  | 0 |  |  | 0 |  |  | 0
Power supply | 1 |  |  | 0 |  |  | 0 |  |  | 0
|  |  |  | 0 |  |  | 0 |  |  | 0
|  |  |  | 0 |  |  | 0 |  |  | 0
|  |  |  | 0 |  |  | 0 |  |  | 0
|  |  |  | 0 |  |  | 0 |  |  | 0
|  |  |  | 0 |  |  | 0 |  |  | 0
|  |  |  | 0 |  |  | 40 |  |  | 780
Note on N6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: cheap easy to use controller. should have enough I/O. would need a display screen and maybe buzzer
Note on N8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: stanard LPG regulators work around the pressure range which we are working with (reducing ~5bar supply pressure down to ~40mbar inspritory pressure)
Note on N9: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: most of these solenoids are 230v...
Note on N10: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: car o2 meter may not be suitable due to it needing hot exhaust fumes to heat the sensor element properly. new ones have their own internal heater but this may not be enough on its own
Note on N11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: cheap common filter hoising. can get the HEPA filter material in bulk or it is used in some vacuum cleaner filters.